## Slide 1
Rachel Escue Helms, PhD, Microscopy Manager
Advanced Imaging Core (AIC)
Elyra 7: Zen Black Introduction and Light Path Settings

## Slide 2: ZEN BLACK INTRO
Zen Black is the software program used for acquiring images on the Elyra 7. The program also performs post-acquisition processing and method-specific processing, including for structured illumination microscopy (SIM) and single-molecule localization microscopy (SMLM) images.
Files are saved as .czi (the proprietary Carl Zeiss Image format) to preserve image metadata, but they can be exported in all major image file formats such as .tif, .png, and .jpg.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 3: ZEN BLACK INTRO
Zen Black functions are separated into 4 tabs:
  LOCATE, ACQUISITION, PROCESSING, and MAINTAIN.
  You will only be using the first three tabs (if anything on the MAINTAIN tab needs to be adjusted, please inform the Core Manager.)
  The most commonly used/altered settings will be covered in this guide.
    For complete settings, refer to the Zen Black manual (available in the AIC).
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 4: OPENING THE PROGRAM
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 5: OPENING THE PROGRAM
Select START SYSTEM in the pop-up.
1
2
An error log automatically begins when the system does not register the incubation and heating systems. This is normal, however, please report any additional errors to the Core Manager.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 6: OPENING THE PROGRAM
Select VIEW > SHOW ALL (GLOBAL)
  Many settings are not automatically visible when first opening the program without turning this option on.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 7: LOCATE TAB
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 8: LOCATE TAB
This tab is for scanning the sample through the eye piece.
  Three pre-set light path buttons are available for your convenience.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 9: LOCATE TAB
The TL (TRANSMITTED LIGHT) option is for brightfield viewing of the sample.
Turn this up to 100% via the lightbulb icon for a lamp-like illumination within the laser-safe box.
The FS77 option uses the FSet77HE filter, allowing you to safely view a fluorescence illumination of the sample.
Choose the ALL OFF option after finding your desired area and before proceeding to the ACQUISITION tab.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 10: LOCATE TAB – OBJECTIVE SELECTION
The objective can be changed by clicking on the OBJECTIVE icon and choosing from the dropdown menu
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 11: ACQUISITION TAB
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 12: ACQUISITION TAB
Note: The bulk of your imaging time will be spent on this tab.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 13: ACQUISITION TAB
One of the easiest ways to set up an experiment is to open an image from your last successful experiment and “reuse” the settings by clicking the RECYCLE button.
  NOTE: This is a shortcut and should not be substituted for confirming your settings before each experiment.
  NOTE: Objective selection is independent from this option, and reusing settings will not make the system bring up the previous objective.
Independently selecting the light-path settings without a previous image will be covered next.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 14: ACQUISITION TAB – ACQUISITION MODE
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 15: ACQUISITION TAB – ACQUISITION MODE
Click the X*Y button to open a dropdown list of common image sizes.
The SCAN AREA shows where within the available field of view you are imaging (most useful when cropping).
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 16: LIGHT PATH SETTINGS
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 17: LIGHT PATH SETTINGS
1 IMAGING SETUP box
  Selection of imaging modality (Widefield, Apotome, Lattice SIM, Laser WF)
  Displays light path for selected modality, including grating, filter sets, and beam splitters.
2 CHANNELS box
  Track and channel specific settings, including laser power and exposure times
1
2
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 18: LIGHT PATH SETTINGS – CHANNELS
When working with live cells, using the dual cameras will decrease the imaging time.
When working with fixed cells/tissues and time is not a factor, keeping the channels on separate tracks is often more convenient.
Camera selections can be made in the adjacent IMAGE SETUP box (covered next).
NOTE:
The terms “tracks” and “channels” are often used interchangeably when talking about image acquisition settings. In the case of the Elyra, there are dual cameras, allowing for 2 channels to be imaged on 1 track.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 19: LIGHT PATH SETTINGS – CHANNELS
To add a new track
1) Click the “+”
A new “blank” track will be listed
2) Name the track
Not an essential step, but future-you will be grateful. Track info (including user-defined names) are listed in the image metadata.
3) Select the lasers and appropriate laser power
If it is a new experiment/fluorophore, start @ 1-1.5% laser power and then increase (or decrease) as necessary.
4) Set the exposure time
A new “blank” track will be listed
3) Select color you want displayed for each channel
Clicking the down arrow opens a pop-up for color selection.
Saving the track configuration is not recommended, as there’s no guarantee another user will not modify it. If you have a configuration that works for you, open your images and recycle the settings.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 20: LIGHT PATH SETTINGS – CHANNELS
Selecting LASER WF as the imaging mode opens up additional settings in the FOV section related to the 3 different illumination modes (Epi/HiLo/TIRF).
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 21: LIGHT PATH SETTINGS – IMAGING SETUP
While settings can be adjusted to user specifications, Zen Black has many default settings for the light path depending on imaging modality and laser line being used (for example, the size grating used for SIM imaging)
User input is mainly needed for the selection of the main beam splitter (light reflecting onto and exciting the sample) and the secondary beam splitter (splitting the signal between the two cameras).
For all experiments requiring laser illumination, using the LBF 405/488/561/642 filter is recommended. (The other filter sets are better suited for widefield/ brightfield illumination.)
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 22: LIGHT PATH SETTINGS – IMAGING SETUP
The secondary beam splitter routes the signal to the camera(s).
  SBS LP 560 – allows light above 560nm to pass through to TV 1, and anything below 560nm is directed to TV 2.
  SBS BP 490-560 + LP 640 – allows light between 490nm and 560 nm or light above 640nm to pass to TV 1, and anything outside of this range to pass to TV 2.
Checking the box under TV 1 and/or TV 2 sets up that camera to capture images for the selected track.
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 23: BONUS! HOW DO I KNOW WHEN I HAVE SET THE LASER TO AN APPROPRIATE LEVEL?
If it is your first time imaging on the Elyra, or first time using a specific fluorophore, you will need to determine the laser attenuation (power) percentage for each channel individually.
Start by getting an image in focus using the LIVE setting on the ACQUISITION tab (we’ll go into more detail on this later)
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS

## Slide 24
Not too bad, but it looks like some areas are saturated (therefore, structures are not well defined)
Keeping it in LIVE mode, let’s change the display from green to “Range Indicator” (dropdown box in CHANNELS)

## Slide 25
The range indicator mode let’s you know if your image has areas that are too bright (shown in red) or too dim (shown in blue, but none in this image).
For this image, we can see that we need to turn the laser down a bit.

## Slide 26
Much better!

## Slide 27
Much better! Notice the structural elements revealed after decreasing the laser. While this will be subjective to your experimental targets, you will generally want to value clarity over intensity.
BEFORE
AFTER

## Slide 28
You are now ready to take an image!
Don’t worry, we’ll walk you through that too. 😉
ADVANCED IMAGING CORE
ELYRA 7: ZEN BLACK INTRODUCTION AND LIGHT PATH SETTINGS